POLICY

<<Facility Name>> wants to ensure the safest possible daily operation and functions of the facility with magnetic resonance imaging (MRI) services to avoid death, injury, or illness.  This policy will guide the testing criteria needed to ensure the safe operation of the MRI equipment.

PURPOSE

To perform annual performance evaluations on MRI equipment.

PROCEDURE

- At least annually, a diagnostic medical physicist or MRI Scientist conducts a performance evaluation, which is documented in Attachment I.

- The evaluation results, along with recommendations for correcting any problems identified, are documented in the Maintenance Management System (MMS).

- The evaluation includes the use of phantoms to assess the following imaging metrics:

  - Image uniformity for all radiofrequency (RF) coils used clinically

  - Signal-to-noise ratio (SNR) for all coils used clinically

  - Slice thickness accuracy

  - Slice position accuracy

  - Alignment light accuracy

  - High-contrast resolution

  - Low-contrast resolution (or contrast-to-noise ratio)

  - Geometric or distance accuracy

  - Magnetic field homogeneity

  - Artifact evaluation

- The performance evaluation is completed by <<DATE>> each year. The information is given to the <<Director of Radiology or Medical Director>> immediately if problems are identified. If information is meeting the criteria determined to be safe, the report will be given within <<24 hours>>.

- The <<Director of Radiology or Medical Director>> will present this report’s information and status to the <<Environment of Care Committee, Radiology Committee, and/or Safety Committee>> at the next meeting after the report.

Attachments:

Attachment I- MRI Services Performance Evaluation (ME-EC.02.04.03EP22b)